Date: 26.09.2024

Urgent Field Safety Notice (RECALL)

10mm FLEXTUBE RESUS BREATHING SYSTEMS FOR USE WITH NEOPUFF® RESUSCITATORS WITH VARIABLE PEEP

For Attention of*: MDSO’s, All clinical staff, Managers and users of the above products

Contact details of local representative (name, e-mail, telephone, address etc.)*

Giedrius Budrys Customer Resolution and Relationship Manager Intersurgical UAB Arnioniu str 60, LT-18170 Pabrade Lithuania Email: giedriusb@intersurgical.lt Tel. +370 387 66611 Fax: +370 387 66622 or This could be a distributor or local branch of the manufacturer. To be added at the appropriate stage in the different local languages

Urgent Field Safety Notice (RECALL)

10mm FLEXTUBE RESUS BREATHING SYSTEMS FOR USE WITH NEOPUFF® RESUSCITATORS WITH VARIABLE PEEP

Risk addressed by FSN

| - Information on Affected Devices* |  |
| --- | --- |
| 1. | - Device Type(s)* |
|  | 10mm Flextube Resus Breathing Systems for use with Neopuff® Resuscitators with Variable PEEP |
| 1. | - Commercial name(s) |
|  | 10mm Flextube neonatal resuscitation breathing system with variable PEEP, double swivel elbow, Neopuff® and universal connectors. ≥ 2.0m 10mm Flextube neonatal resuscitation breathing system with variable PEEP, double swivel elbow for Neopuff®, ≥ 1.2m 10mm Flextube neonatal resuscitation breathing system with variable PEEP, double swivel elbow for Neopuff®, ≥ 2m 10mm Flextube neonatal resuscitation breathing system for use with Neopuff® resuscitators with variable PEEP and double swivel elbow. ≥ 2.4m 10mm Flextube neonatal resuscitation breathing system with variable PEEP, double swivel elbow, Neopuff® and universal connectors, ≥ 1.2m |
| 1. | - Unique Device Identifier(s) (UDI-DI) |
|  | 5030267079377 5030267087310 5030267090440 5030267095773 5030267105038 |
|  | - Primary clinical purpose of device(s)* |
|  | The intended use of this product is to deliver and remove gases to and from a patient in order to resuscitate/ventilate a neonatal patient. |
| 1. | - Device Model/Catalogue/part number(s)* REF: 6431002 REF: 6432000 REF: 6432001 REF: 6432002 REF: 6433000 |
| 1. | - Software version |
|  | N/A |
| 1. | - Affected serial or lot number range sold worldwide: 32400028 32400306 32400567 32400797 32401191 32401580 32401900 32401903 32401971 32402073 32402162 32402626 32403853 32404419 32404497 32404565 32405021 32405255 32406256 32406258 32406864 32407341 32407377 32407649 32407909 32408160 32408388 32408457 32408788 32410623 |
| 1. | - Associated devices |
|  | N/A. |

| - Reason for Field Safety Corrective Action (FSCA)* |  |  |  |
| --- | --- | --- | --- |
| 2. | - Description of the product problem* |  |  |
|  | With some systems it is not possible to achieve a secure connection of the pink Neopuff® connector to the Neopuff® Resuscitator as shown below. |  |  |
| 2. | - Hazard giving rise to the FSCA* |  |  |
|  | A loose connection between the breathing system and the resuscitation device could result in delay to treatment or a disconnection during use. The intended use of this product is to deliver and remove gases to and from a patient in order to resuscitate/ventilate a neonate, a delay or disconnection would prolong/induce asphyxia, potentially resulting in patient brain damage or death. |  |  |
| 2. | 3. Probability of problem arising |  |  |
|  | High in the affected Lot number range, as a large number of pink Neopuff® connectors are potentially affected. |  |  |
| 2. | - Predicted risk to patient/users |  |  |
|  | The risk of patient harm has been evaluated as major however, the probability of occurrence of patient harm has been assessed as rare. The pink Neopuff® connector is supplied as an accessory for some of the affected products, and it is therefore not always used. Where the pink Neopuff® connector is used, any loose connection will be identified when attaching to the Neopuff® Resuscitator. We believe it is essential to address the issue promptly to further reduce the risk of any potential patient harm. |  |  |
| 2. | - Further information to help characterise the problem |  |  |
|  | N/A |  |  |
| 2. | - Background on Issue |  |  |
|  | Following a customer report from the market and subsequent thorough inspection and analysis of returned samples and internal stock, we have identified a potential safety concern related to the Resuscitation Breathing Systems for use with the Neopuff Resuscitator, as listed above. Unfortunately some products have been manufactured where the female taper of the pink Neopuff® connector is oversized, which could result in an insecure connection to the Neopuff® Resuscitator. This problem only relates to the products and Lots listed above, which have all been manufactured and supplied this year. |  |  |
| 2. | - Other information relevant to FSCA N/A |  |  |
|  | - Type of Action to mitigate the risk* |  |  |  |
| 3. | - Action To Be Taken by the User* ☒ Identify Device ☒ Quarantine Device ☐ Return Device ☒ Destroy Device ☐ On-site device modification/inspection ☐ Follow patient management recommendations ☐ Take note of amendment/reinforcement of Instructions For Use (IFU) ☐ Other ☐ None Identify and immediately quarantine all affected codes and lot numbers listed above and do not use these devices. Please complete the Reply Form to confirm the products have been disposed of locally or to arrange collection of the devices and a credit. If you have no affected devices in stock, please confirm this using the Reply Form. Return the completed Reply Form to giedriusb@intersurgical.lt . Please continue to report to Intersurgical any adverse events involving this product. |  |  |  |
| 3. | - By when should the action be completed? | Immediately on receipt of this FSN and ongoing until no affected stock listed in this FSN is remaining. |  |  |
| 3. | - Particular considerations for: N/A Is follow-up of patients or review of patients’ previous results recommended? Not applicable. |  |  |  |
| 3. | - Is customer Reply Required? * (If yes, form attached specifying deadline for return) |  | Yes |  |
| 3. | - Action Being Taken by the Manufacturer ☒ Product Removal ☐ On-site device modification/inspection ☐ Software upgrade ☐ IFU or labelling change ☐ Other ☐ None |  |  |  |
| 3 | - By when should the action be completed? | One month from receipt of the FSN |  |  |
| 3. | - Is the FSN required to be communicated to the patient /lay user? |  | No |  |
| 3 | - If yes, has manufacturer provided additional information suitable for the patient/lay user in a patient/lay or non-professional user information letter/sheet? |  |  |  |
|  | N/A |  |  |  |

|  | - General Information* |  |
| --- | --- | --- |
| 4. | - FSN Type* | New – Recall |
| 4. | - For updated FSN, reference number and date of previous FSN | N/A |
| 4. | - For Updated FSN, key new information as follows: |  |
|  | N/A |  |
| 4. | - Further advice or information already expected in follow-up FSN? * | No |
| 4 | - If follow-up FSN expected, what is the further advice expected to relate to: |  |
|  | N/A |  |
| 4 | - Anticipated timescale for follow-up FSN | N/A |
| 4. | - Manufacturer information (For contact details of local representative refer to page 1 of this FSN) |  |
|  | - Company Name | Intersurgical Ltd. |
|  | - Address | Crane House, Molly Millars Lane, Wokingham, Berkshire, RG41 2RZ |
|  | - Website address | https://www.intersurgical.com/ |
| 4. | - The Competent (Regulatory) Authority of your country has been informed about this communication to customers. * |  |
| 4. | - List of attachments/appendices: | FSCA, Distributor/Customer Reply Form |
| 4. | - Name/Signature | Ivan Seniut, Group Quality and Regulatory Affairs Director, Intersurgical |
|  | Transmission of this Field Safety Notice |  |
|  | This notice needs to be passed on all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. (As appropriate) Please transfer this notice to other organisations on which this action has an impact. (As appropriate) Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. |  |