Date: 2021-08-11

Urgent Field Safety Notice

Central Monitoring System MFM-CMS V2.66

For Attention of*:The detailed serial number and other tracking information, please see the attachment for < List of customer>

Contact details of local representative (name, e-mail, telephone, address etc.)*

EDAN Instruments GmbH, Monday to Friday 09:00-17:00 (UTC +01:00)Tel: +49 (0) 6103 202 0781

Urgent Field Safety Notice (FSN)

Central Monitoring System MFM-CMS V2.66

Risk addressed by FSN

| - Information on Affected Devices* |  |
| --- | --- |
| 1. | - Device Type(s)* |
|  | Central Monitoring System MFM-CMS V2.66 |
| 1. | - Commercial name(s) |
|  | Central Monitoring System |
| 1. | - Unique Device Identifier(s) (UDI-DI) |
|  | 06944413800229 |
| 1. | - Primary clinical purpose of device(s)* |
|  | MFM-CMS provides centralized monitoring and critical care management for patients monitored by EDAN bedside monitors. From MFM-CMS, clinicians can gain access to patient information for patients on the Network. MFM-CMS displays waveforms, parameters and alarm status of EDAN bedside monitors for up to 32 patients on a single screen or up to 64 patients using two screens. |
| 1. | - Device Model/Catalogue/part number(s)* |
|  | MFM-CMS |
| 1. | - Software version |
|  | V2.66 |

| - Reason for Field Safety Corrective Action (FSCA)* |  |
| --- | --- |
| 2. | - Description of the product problem* |
|  | Recently during internal testing, it was found that under specific circumstances, MFM-CMS Central Monitoring System V2.66 version did not effectively display SpO2 alarm information. |
| 2. | - Hazard giving rise to the FSCA* |
|  | When healthcare professionals only pay attention to the alarm on the MFM-CMS, it may cause patients with SpO2 below the alarm limit unable to get timely attention. Edan has not received any report of serious incident that caused patient injury regarding this problem. |

|  | - Type of Action to mitigate the risk* |  |  |
| --- | --- | --- | --- |
| 3. | - Action To Be Taken by the User* ☐ Identify Device ☐ Quarantine Device ☐ Return Device ☒ Destroy Device ☒ On-site device modification/inspection ☐ Follow patient management recommendations ☐ Take note of amendment/reinforcement of Instructions For Use (IFU) ☐ Other ☐ None Provide further details of the action(s) identified. |  |  |
| 3. | - By when should the action be completed? |  |  |
| 3. | - Particular considerations for: Choose an item. Is follow-up of patients or review of patients’ previous results recommended? Choose an item. Provide further details of patient-level follow-up if required or a justification why none is required |  |  |
| 3. | - Is customer Reply Required? * (If yes, form attached specifying deadline for return) |  | Yes |
| 3. | - Action Being Taken by the Manufacturer ☐ Product Removal ☐ On-site device modification/inspection ☒ Software upgrade ☒ IFU or labelling change ☐ Other ☐ None Provide further details of the action(s) identified. |  |  |
| 3 | - By when should the action be completed? | November 13, 2021 |  |
| 3. | - Is the FSN required to be communicated to the patient /lay user? |  | No |
| 3 | - If yes, has manufacturer provided additional information suitable for the patient/lay user in a patient/lay or non-professional user information letter/sheet? |  |  |
|  | No Choose an item. |  |  |

|  | - General Information* |  |
| --- | --- | --- |
| 4. | - FSN Type* | New |
| 4. | - For updated FSN, reference number and date of previous FSN | Provide reference and date of previous FSN if relevant |
| 4. | - For Updated FSN, key new information as follows: |  |
|  | Summarise any key difference in devices affected and/or action to be taken. |  |
| 4. | - Further advice or information already expected in follow-up FSN? * | No |
| 4 | - If follow-up FSN expected, what is the further advice expected to relate to: |  |
|  | Eg patient management, device modifications etc |  |
| 4 | - Anticipated timescale for follow-up FSN | For provision of updated advice. |
| 4. | - Manufacturer information (For contact details of local representative refer to page 1 of this FSN) |  |
|  | - Company Name | Edan Instruments, Inc. |
|  | - Address | #15 Jinhui Road,Jinsha Community, Kengzi Sub-District, Pingshan District, 518122 Shenzhen, P.R.China. |
|  | - Website address | https://www.edan.com/ |
| 4. | - The Competent (Regulatory) Authority of your country has been informed about this communication to customers. * None |  |
|  | Transmission of this Field Safety Notice |  |
|  | 1. Edan sends field safety notice to affected customers, requesting them to stop using V2.66 MFM-CMS. 2. Edan will upgrade the MFM-CMS for affected customers. 3. Other versions of MFM-CMS work properly. They are not affected and can be used without any problem. |  |

Note: Fields indicated by * are considered necessary for all FSNs. Others are optional.