URGENT FIELD SAFETY NOTICE/ DEVICE RECALL

20/30 INDEFLATORTM, INDEFLATORTM Plus 30 and Priority Packs

Commercial Name:

20/30 INDEFLATORTM Inflation Device

20/30 Priority Pack Accessory Kit

INDEFLATORTM PLUS 30

Plus 30 Priority Pack Accessory Kit

FSCA-Identifier:  INDEFLATORS & PPacks March 11, 2022

Manufacturer: Abbott Vascular Santa Clara SRN# US-MF-000003850

Type of Action:  Device Recall

## Attention:  Risk Manager or Healthcare Professional

Dear Valued Abbott Customer:

Abbott has initiated a field action for specific lots of 20/30 INDEFLATORTM, INDEFLATORTM Plus 30 and associated Priority Packs. Our records indicate that affected devices have been shipped to your account.

This action does not affect patients having successfully undergone procedures using these devices.

Devices from these lots may exhibit leaks and/or a loose connection at the rotating luer assembly or stopcock connection, which could lead to air ingress under vacuum. Analysis indicates an estimated rate of occurrence of air ingress associated with the device is 0.4%.  While no long-term adverse patient effects have been attributed to this issue, potential risks include air embolism, thrombosis and foreign body in patient.

What action should you to take?

- Immediately stop using devices from affected lots (see attached)

- Review your inventory, complete and return the provided Effectiveness Check Form

- Return all unused affected devices to Abbott

- Share this notification with relevant personnel in your organization

- Report any occurrence of product performance issues or patient adverse events to Abbott

What action is Abbott taking?

- Abbott has taken immediate action to stop shipping devices from affected lots.

- The investigation has determined there are no other affected products or lots in distribution.

- Abbott will implement appropriate corrective actions to ensure product performance.

- Abbott will work with you to replace inventory, when available.

- The appropriate regulatory agencies have been notified of this action.

We regret any inconvenience this may cause you and appreciate your patience.  Abbott is committed to providing high quality, compliant products and ensuring customer satisfaction.  If you have any questions, please do not hesitate to contact your local Abbott Representative or Customer Service department.

Sincerely,

<signature of country manager>

<printed name>

Part Numbers and Lot Numbers

| Device Identifier/GTIN | Device Description | Part Number | Lot Number |  |
| --- | --- | --- | --- | --- |
| 08717648013591 | 20/30 INDEFLATOR | 1000184 | 60309678 | 60315914 | 60331726 |
|  |  |  | 60311336 | 60317535 | 60331727 |
|  |  |  | 60311337 | 60317539 | 60334491 |
|  |  |  | 60311338 | 60318209 | 60334492 |
|  |  |  | 60311339 | 60320071 | 60337162 |
|  |  |  | 60315913 | 60320899 | 60337166 |
| 08717648013614 | 20/30 Priority Pack Accessory Kit/.096 RHV | 1000186 | 60317536 | 60320910 | 60334116 |
|  |  |  | 60317537 | 60320911 | 60338482 |
|  |  |  | 60317542 | 60325409 | 60338488 |
|  |  |  | 60318666 | 60326622 | 60338489 |
|  |  |  | 60320079 | 60326623 |  |
|  |  |  | 60320909 | 60329936 |  |
| 08717648015274 | 20/30 Priority Pack Kit/.115 RHV | 1000186-115 | 60311340 | 60318663 | 60329334 |
|  |  |  | 60311346 | 60318664 | 60334117 |
|  |  |  | 60317538 | 60318665 | 60334737 |
|  |  |  | 60318661 | 60318667 |  |
|  |  |  | 60318662 | 60325410 |  |
| 08717648013973 | 20/30 Priority Pack Accessory Kit w/Copilot | 1003327 | 60308571 | 60316407 | 60325101 |
|  |  |  | 60308572 | 60317004 | 60325103 |
|  |  |  | 60308573 | 60317279 | 60326298 |
|  |  |  | 60308574 | 60317280 | 60326299 |
|  |  |  | 60308575 | 60317533 | 60326300 |
|  |  |  | 60309671 | 60317540 | 60326301 |
|  |  |  | 60309672 | 60317541 | 60326425 |
|  |  |  | 60309673 | 60317947 | 60326859 |
|  |  |  | 60309674 | 60317948 | 60326860 |
|  |  |  | 60309675 | 60318668 | 60326861 |
|  |  |  | 60309676 | 60318669 | 60326862 |
|  |  |  | 60309677 | 60318670 | 60326863 |
|  |  |  | 60309681 | 60319819 | 60328011 |
|  |  |  | 60309682 | 60320067 | 60328023 |
|  |  |  | 60309683 | 60320068 | 60328355 |
|  |  |  | 60309684 | 60320069 | 60328356 |
|  |  |  | 60309685 | 60320070 | 60329330 |
|  |  |  | 60309686 | 60320072 | 60329331 |
|  |  |  | 60309687 | 60320073 | 60329332 |
|  |  |  | 60311341 | 60320074 | 60329333 |
|  |  |  |  | 60320075 | 60329967 |
| 08717648013973 | 20/30 Priority Pack Accessory Kit w/Copilot | 1003327 | 60311342 | 60320076 | 60329968 |
|  |  |  | 60311343 | 60320914 | 60329969 |
|  |  |  | 60311344 | 60322147 | 60330058 |
|  |  |  | 60311345 | 60322182 | 60331041 |
|  |  |  | 60312167 | 60322183 | 60331042 |
|  |  |  | 60312168 | 60322184 | 60331043 |
|  |  |  | 60312169 | 60322185 | 60331358 |
|  |  |  | 60312170 | 60323316 | 60331537 |
|  |  |  | 60312171 | 60323317 | 60331538 |
|  |  |  | 60312172 | 60323318 | 60331731 |
|  |  |  | 60312173 | 60323319 | 60331733 |
|  |  |  | 60312174 | 60323320 | 60331943 |
|  |  |  | 60312175 | 60323321 | 60334120 |
|  |  |  | 60312176 | 60323322 | 60334121 |
|  |  |  | 60312177 | 60323323 | 60334122 |
|  |  |  | 60312178 | 60323324 | 60334123 |
|  |  |  | 60313415 | 60323325 | 60334449 |
|  |  |  | 60313416 | 60323434 | 60335132 |
|  |  |  | 60313417 | 60323785 | 60335214 |
|  |  |  | 60313418 | 60323786 | 60335817 |
|  |  |  | 60313420 | 60323787 | 60336487 |
|  |  |  | 60313421 | 60325097 | 60337158 |
|  |  |  | 60313422 | 60325098 | 60337397 |
|  |  |  | 60315918 | 60325099 |  |
|  |  |  | 60315919 | 60325100 |  |
| 08717648013584 | PLUS 30 INDEFLATOR | 1000183 | 60312165 | 60320908 | 60327731 |
|  |  |  | 60313282 | 60320913 | 60328010 |
|  |  |  | 60313413 | 60322176 | 60328352 |
|  |  |  | 60315909 | 60322179 | 60333403 |
|  |  |  | 60316775 | 60326322 | 60337153 |
|  |  |  | 60320904 | 60326865 | 60337157 |
|  |  |  | 60320905 | 60326866 | 60338458 |
|  |  |  | 60320906 | 60326867 | 60338479 |
| 8717648013607 | Plus 30 Priority Pack 0.096 | 1000185 | 60312166 | 60317532 | 60326601 |
|  |  |  | 60313414 | 60323313 | 60329935 |
|  |  |  | 60315910 | 60323314 | 60337154 |
| 08717648015267 | Plus 30 Priority Pack 0.115 | 1000185-115 | 60316763 | 60323315 |  |

Commercial Name:

20/30 INDEFLATORTM Inflation Device

20/30 Priority Pack Accessory Kit

INDEFLATORTM PLUS 30

Plus 30 Priority Pack Accessory Kit

FSCA-Identifier:  INDEFLATORS & PPacks March 11, 2022

Manufacturer: Abbott Vascular Santa Clara, SRN# US-MF-000003850

Type of Action:  Device Recall

Effectiveness Check Form

Customer Account #	________________________________________

Account Name	________________________________________

Address		________________________________________

________________________________________

(Information required for regulatory effectiveness check)

After reviewing your inventory for the affected devices, complete this form and return this form and any affected devices to Abbott per the instructions below.

Check One:

A thorough search for all affected devices has been completed and no affected units remain in inventory. No devices will be returned.

Affected devices have been identified and are being returned

## RGA Number: __________________________

____________________________	___________________________            __________

##### Customer Name/ Job Title (print)         Signature 			                   Date

This form is to be returned to Abbott

- If returning product, call Customer Service <insert phone number> to receive RGA number. Record RGA number above.

Scan and email this form to <insert email> or fax to <insert phone number>

- Return a copy of this completed form with the returned product.