IMPORTANT:

URGENT FIELD SAFETY NOTICE

BIOFIRE® FILMARRAY® TORCH System – Ref Number(s) HTFA-ASY-0104, HTFA-ASY-0001

BIOFIRE® FILMARRAY® TORCH System Power Supply Correction

Please distribute the attached customer letter.

To the Laboratory Manager

To the attention of the Laboratory Medical Director

Date

bMx local contact information

(to be adapted at local level)

Our reference: FSCA - FIELD SAFETY CORRECTIVE ACTION - FA-TWD-000081 (FSCA 5761 Correction)

| Impacted products (to be adapted at local level if necessary including for names and ref #, local license #, name and address of manufacturer) |  |  |  |
| --- | --- | --- | --- |
| Product Name | Reference Number | Lot Number/Serial Number/ Product version | Product Expiration Date (if applicable) |
| BIOFIRE® FILMARRAY® TORCH System Base | HTFA-ASY 0104, Shipping Configuration No. HTFA ASY-0001 | Manufactured before September 16, 2024 | N/A |

Dear bioMérieux Customer,

Our records indicate that you may be using the affected product listed in the table above.

The purpose of this letter is to inform you of a field deployable solution to resolve the previous field safety notice FSCA 5761 initially communicated in October 2023, involving the BIOFIRE® FILMARRAY® TORCH System Base (reference number: HTFA-ASY-0104, HTFA-ASY-0001) and an increased risk of degradation of the power switch of the BIOFIRE FILMARRAY TORCH Base.

Required actions In this context, we request you to take the following actions: - Please install, according to Installation procedure in Appendix A, the power cord with fuse (Field replacement unit 6209340 Harness, Fuse Extension Cable) upon receipt from bioMérieux. [OR] - bioMérieux will contact you to coordinate a date and send a field service representative to install the power cord with fuse at site. - Distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product. - If you have further distributed this product, please identify any recipients and notify them at once. - Complete the Acknowledgement Form in Attachment A and return it to your local bioMérieux representative (to be adapted at local level) to confirm receipt of this notice. It is important that you return the acknowledgement form to bioMérieux even if you determine that this urgent product correction notice does not impact your facility.

Description of the issue

Our CAPA investigation determined this increased risk may be seen on BIOFIRE FILMARRAY TORCH systems that are power cycled frequently. Power cycling can lead to arcing inside of the power switch and may result in carbon build-up and/or deformation of the power switch case. Carbon build-up and/or deformation may result in an open circuit causing the power switch to fail. If carbon build-up and/or deformation of the power switch has already occurred, regardless of power cycle frequency, use with multiple modules over time increases the risk of failure.

As of September 16, 2024, a power entry module has been redesigned for new production of BIOFIRE FILMARRAY TORCH System Bases to prevent the risk of deformation or degradation and subsequent excess heat or electrical short. For BIOFIRE FILMARRAY TORCH System Bases produced before September 16, 2024, bioMérieux has developed a fused extension cable which should be installed on each BIOFIRE FILMARRAY TORCH System Base.

Impact to User/Customer/Patients

The fuse between the power source and power switch would open the circuit and prevent risk of fire, burns, smoke, and electrical shock to the operator, at the power entry module. The fused extension cable will not prevent or resolve existing carbon build-up and/or deformation, thus in addition to use of the fused extension cable, bioMérieux still recommends that BIOFIRE FILMARRAY TORCH systems with 6 or more modules have a power switch replacement, as indicated in previous letters. Please continue to operate the instrument according to the instructions of the Operator Manual.

Local legal mentions to be added if necessary at local level (e.g. in case of recall, reporting to NCA, recall methods)

bioMérieux is committed to providing our customers with the highest quality product possible.

We sincerely apologize for any inconvenience that this may have caused you. If you require additional assistance or have any questions, please contact your local bioMérieux Customer Service representative (to be adapted at local level).

Kind Regards,

Customer Service

Attachment A: Acknowledgement Form.

Urgent: Field Safety Notice

FSCA - FIELD SAFETY CORRECTIVE ACTION FA-TWD-000081

BIOFIRE® FILMARRAY® TORCH System Power Supply Correction

##### To be returned to your bioMerieux customer service (to be adapted at local level)  at the following

##### FAX number: XXXXXXXX OR EMAIL ADDRESS: XXXXXXX

| Name and Address of the laboratory |  |
| --- | --- |
| Contact information |  |
| Customer Account Number |  |

Local legal mentions to be added if necessary at local level)

- ☐ I am not impacted by the issue. Please provide rationale: ……………………………………

- ☐ I have implemented the required actions.

- Table to be added and adapted if necessary to monitor quantities received/discarded (products names and ref.# to be adapted at local level if necessary) depending on the required actions.

| REF # | Product Name | Serial number #/ Version number | Quantity impacted | Quantity corrected* |
| --- | --- | --- | --- | --- |
| HTFA-ASY 0104, Shipping Configuration No.: HTFA ASY-0001 | BIOFIRE® FILMARRAY® TORCH | N/A |  |  |

bioMerieux subsidiary or distributor 
* Quantity corrected include installation of fused extension cable

If you have encountered impact on patients’ results, or reports of illness or injury related to the identified issue and have not yet reported to bioMérieux, please contact your local bioMérieux representative.

date………………	signature……………………………………………

It is important that you complete this Acknowledgement Form and return it to bioMérieux