IMPORTANT:

URGENT FIELD SAFETY NOTICE

False Resistant Antibiotic results (cza02n, ipm05n, ipr01n, mev01n, pb02n) in VITEK® 2 AST GN cards for Enterobacterales species and/or Pseudomonas aeruginosa

Please distribute the attached customer letter.

To the Laboratory Manager

To the attention of the Laboratory Medical Director

Date

bMx local contact information

(to be adapted at local level)

Our reference: FA-TWD-000064- FSCA - FIELD SAFETY CORRECTIVE ACTION

Dear bioMérieux Customer,

This communication is to inform you of a potential risk of Quality Control failures and/or false resistant antibiotic results when testing isolates of Enterobacterales/Enterobacteriaceae species and/or Pseudomonas aeruginosa with the following antibiotic formulations: Ceftazidime/Avibactam (cza02n), Imipenem (ipm05n), Imipenem/Relebactam (ipr01n), Meropenem/Vaborbactam (mev01n), and Polymyxin B (pb02n).

Our records indicate that your laboratory received one or more products listed in table 1 below.

Table 1: List of impacted cards

| Product Name | Reference | Drug Impacted |
| --- | --- | --- |
| VITEK® 2 AST-N399 | 423701 | ipm05n |
| VITEK® 2 AST-N402 | 423644 | ipm05n |
| VITEK® 2 AST-N403 | 423645 | cza02n/ipm05n |
| VITEK® 2 AST-N404 | 423664 | ipm05n |
| VITEK® 2 AST-N405 | 423864 | ipm05n |
| VITEK® 2 AST-N406 | 423869 | ipm05n |
| VITEK® 2 AST-N407 | 423870 | cza02n |
| VITEK® 2 AST-N409 | 423925 | cza02n |
| VITEK® 2 AST-N412 | 423936 | ipm05n |
| VITEK® 2 AST-N413 | 423928 | ipm05n |
| VITEK® 2 AST-N414 | 423933 | ipm05n |
| VITEK® 2 AST-N415 | 423934 | ipm05n |
| VITEK® 2 AST-N416 | 423935 | ipm05n |
| VITEK® 2 AST-N417 | 423880 | ipm05n |
| VITEK® 2 AST-N419 | 423948 | cza02n/ipm05n |
| VITEK® 2 AST-N421 | 424055 | cza02n/ipm05n |
| VITEK® 2 AST-N428 | 424320 | ipm05n |
| VITEK® 2 AST-N429 | 424321 | ipm05n |
| VITEK® 2 AST-N430 | 424322 | ipm05n |
| VITEK® 2 AST-N433 | 424389 | ipm05n |
| VITEK® 2 AST-N439 | 424501 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-N440 | 424502 | cza02n/ipm05n/ipr01n |
| VITEK® 2 AST-N443 | 424541 | cza02n/ipm05n/ipr01n |
| VITEK® 2 AST-N444 | 424587 | cza02n/ipm05n/ipr01n |
| VITEK® 2 AST-N467 | 424857 | cza02n/ipr01n |
| VITEK® 2 AST-N476 | 424934 | cza02n |
| VITEK® 2 AST-N478 | 425052 | cza02n |
| VITEK® 2 AST-N480 | 425084 | cza02n |
| VITEK® 2 AST-N481 | 425085 | cza02n |
| VITEK® 2 AST-N485 | 425182 | cza02n |
| VITEK® 2 AST-N496 | 425265 | cza02n |
| VITEK® 2 AST-N800 | 423310 | cza02n |
| VITEK® 2 AST-N802 | 423706 | ipm05n |
| VITEK® 2 AST-N804 | 424634 | cza02n/ipr01n |
| VITEK® 2 AST-N808 | 424711 | ipm05n |
| VITEK® 2 AST-N810 | 424712 | cza02n |
| VITEK® 2 AST-N811 | 424713 | cza02n |
| VITEK® 2 AST-N813 | 424722 | cza02n/ipr01n |
| VITEK® 2 AST-N815 | 425054 | cza02n/mev01n |
| VITEK® 2 AST-XN09 | 423425 | cza02n |
| VITEK® 2 AST-XN15 | 423829 | cza02n/mev01n |
| VITEK® 2 AST-XN17 | 423673 | cza02n |
| VITEK® 2 AST-XN18 | 423874 | cza02n |
| VITEK® 2 AST-XN20 | 423947 | cza02n/mev01n |
| VITEK® 2 AST-XN21 | 424197 | cza02n/mev01n |
| VITEK® 2 AST-XN22 | 424199 | cza02n/mev01n |
| VITEK® 2 AST-XN23 | 424198 | caz02n/ipr01n/mev01n/pb02n |
| VITEK® 2 AST-XN24 | 424351 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN26 | 424500 | cza02n/ipm05n/ipr01n/mev01n |
| VITEK® 2 AST-XN28 | 424586 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN29 | 424604 | cza02n/ipr01n |
| VITEK® 2 AST-XN30 | 424639 | cza02n/ipm05n/ipr01n/mev01n/pb02n |
| VITEK® 2 AST-XN31 | 424640 | ipm05n/ipr01n/mev01n/pb02n |
| VITEK® 2 AST-XN32 | 424678 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN33 | 424723 | cza02n/ipm05n/ipr01n/mev01n |
| VITEK® 2 AST-XN34 | 424752 | cza02n/ipr01n/mev01n/pb02n |
| VITEK® 2 AST-XN35 | 424810 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN36 | 424860 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN37 | 424888 | cza02n/ipr01n |
| VITEK® 2 AST-XN38 | 425074 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN39 | 425086 | ipr01n/mev01n |
| VITEK® 2 AST-XN40 | 425095 | cza02n/ipr01n/mev01n |
| VITEK® 2 AST-XN41 | 425204 | cza02n |
| VITEK® 2 AST-XN42 | 425309 | cza02n/ipr01n/mev01n |

Description of the issue

bioMérieux has identified an increase in complaints for out-of-range high (OORH) quality control (QC) results for the following antibiotic formulations: Ceftazidime/Avibactam (cza02n), Imipenem (ipm05n), Imipenem/Relebactam (ipr01n), Meropenem/Vaborbactam (mev01n), and Polymyxin B (pb02n). During the investigation we identified that a small number of VITEK® 2 GN-AST cards users, have also reported false resistant antibiotic results with the same antibiotics when testing strains of Enterobacterales/Enterobacteriaceae species and/or Pseudomonas aeruginosa. Please note that the internal investigation is still ongoing, and the root cause has not yet been fully determined.

Impact to User/Customer/Patients

There are two potential impacts: QC failure and/or false resistant results of patient isolates due to overcalling of the MIC value.  False resistant results may lead to the use of less effective antibiotic treatments or the association of multiple antibiotics having negative side effects, which could potentially impact patient’s health.

Required actions bioMérieux recommends that clinically significant isolates of Enterobacterales/Enterobacteriaceae species and Pseudomonas aeruginosa with resistant results to Ceftazidime/Avibactam (cza02n), Imipenem (ipm05n), Imipenem/Relebactam (ipr01n), Meropenem/Vaborbactam (mev01n), or Polymyxin B (pb02n) be considered for confirmation using an alternative method for all VITEK® 2 GN-AST cards. The determination of clinical significance should be made by each laboratory based on the antimicrobial susceptibility profile and the local availability of alternative antibiotics. If appropriate for your laboratory, create a custom bioART rule (regionally specific as listed in Table 2) when using a VITEK® 2 GN-AST card with the impacted antibiotics. The following are the bioART rule criteria: - If: Organism is [SPECIES] (“Enterobacterales/Enterobacteriaceae” or “P. aeruginosa”) Antibiotic is [DRUG VERSION], Interpretation R (Resistant) - Then: Stop for Review Create a user-defined comment on the lab report to guide the user for further instruction regarding alternate testing (i.e. Per FA-TWD-000064 FSCA, confirm [DRUG NAME] ([DRUG VERSION]) resistance with an alternative method before reporting; only when needed for patient care. - Note: For Enterobacterales/Enterobacteriaceae if software version is 9.03 or prior apply for Enterobacteriaceae, if software version is 9.04 or after apply Enterobacterales. For assistance with bioART rule creation, please contact your local bioMérieux representative. In accordance with your facility’s policies, you may consider performing a retrospective review of clinically significant resistant results. Among tests previously performed, identify any possible false resistant results that may have occurred, analyze the related risks, and determine appropriate actions, if applicable. Please ensure that this information is communicated to all relevant personnel within your laboratory, retain a copy in your records, and forward this information to all parties that may use this product, including those to whom you may have transferred the product. Please complete the Acknowledgement Form in Attachment A and return it to your local bioMérieux representative to confirm receipt of this notice. It is important that you return the acknowledgement form to bioMérieux even if you determine that this product correction notice does not impact your facility. As a reminder, please visit the Resource Center to confirm if a reference package insert is associated with a Field Corrective Action or Field Safety Corrective Action. https://resourcecenter.biomerieux.com/

Table2. Recommended Actions by Region/Country:

| Antibiotic​ | Organism​ | Perform alternate method in case of resistant result for clinically significant isolates?​ (NO, means no bioART rule is needed) ​ |  |  |
| --- | --- | --- | --- | --- |
|  |  | USA​ | EMEA, CANADA and ASPAC​ | LATAM​ |
| Ceftazidime/avibactam (cza02n) | Enterobacterales ​/Enterobacteriaceae*** | YES​ | YES​ | YES​ |
|  | P. aeruginosa​ | NO, Existing limitation*​ | YES​ | YES​ |
| Imipenem (ipm05n) | Enterobacterales ​/ Enterobacteriaceae*** | YES​ | YES​ | YES​ |
|  | P. aeruginosa​ | YES​ | YES​ | YES​ |
| Imipenem/relebactam (ipr01n) | Enterobacterales ​ /Enterobacteriaceae*** | YES​ | YES​ | NO**​ |
|  | P. aeruginosa​ | YES​ | YES​ | NO**​ |
| Meropenem/vaborbactam (mev01n) | Enterobacterales ​/ Enterobacteriaceae*** | YES​ | YES​ | NO**​ |
|  | P. aeruginosa​ | NO, not FDA cleared​ | YES​ | NO**​ |
| Polymyxin B (pb02n) | Enterobacterales ​/ Enterobacteriaceae*** | NO, not FDA cleared​ | NO**​ | NO**​ |
|  | P. aeruginosa​ | YES​ | NO**​ | NO**​ |

* Existing limitation for MIC values of ≥16 µg/mL: resistant results should not be reported

** No cards exist with the impacted drugs, rule to apply if drug is available or new card is created with impacted drugs in future​

*** For Enterobacterales/Enterobacteriaceae if software version is 9.03 or prior apply for Enterobacteriaceae, if software version is 9.04 or after apply Enterobacterales.

Local legal mentions to be added, if necessary, at local level (e.g. in case of recall, reporting to NCA, recall methods)

bioMérieux is committed to providing our customers with the highest quality product possible.

We sincerely apologize for any inconvenience that this may have caused you. If you require additional assistance or have any questions, please contact your local bioMérieux Customer Service representative (to be adapted at local level).

Yours faithfully,

Customer Service

Attachment A: Acknowledgement Form.

Urgent Field Safety Corrective Notice

FSCA - FIELD SAFETY CORRECTIVE ACTION FA-TWD-000064

False Resistant Antibiotic results (ipm05n, pb02n, cza02n, mev01n, ipr01n) in VITEK® 2 AST GN cards for Enterobacterales/Enterobacteriacae species and/or Pseudomonas aeruginosa

##### To be returned to your bioMerieux customer service (to be adapted at local level) at the following

##### FAX number: XXXXXXXX OR EMAIL ADDRESS: XXXXXXX

| Name and Address of the laboratory |  |
| --- | --- |
| Contact information |  |
| Customer Account Number |  |

Local legal mentions to be added, if necessary, at local level)

- ☐ I am not impacted by the issue. Please provide rationale: ……………………………………

- ☐ I am impacted by the issue.

- Impacted References: 	Reference	……………………………………

- Reference	..................................................

- Reference	..................................................

- Other	 	..................................................

- ☐ I have implemented the required actions.

- Table to be added and adapted if necessary to monitor quantities received/discarded (products names and ref.# to be adapted at local level if necessary) depending on the required actions.

Have you encountered an impact on patients’ results, or reports of illness or injury related to the identified issue? (to be completed based on FSCA issue)

☐ Yes    ☐ No

date………………	signature……………………………………………

It is important that you complete this Acknowledgement Form and return it to bioMérieux